Digital Multimeters VOAC7520/7521A Specifications

| * | Accuracy X% of reading ± Y digits indicated by X+Y |
| --- | --- |
|  | Accuracy is assured for 1 year after the calibration is performed. |

| 1. Sample Rate |
| --- |
|  | Sample Rate Resolution Reading Rate Hum Rejection SLOW 5.5-digit approx. 4 times/sec Yes MID 5.5-digit approx. 20 times/sec Yes FAST 4.5-digit approx. 100 times/sec N/A |

| 2. DC Volt (DCV) |
| --- |
|  | Range Resolution Input Resistance Accuracy* 5.5-digit 4.5-digit SLOW/MID FAST 500mV 1µV 10µV 1000M or more 0.012+5 0.012+10 5V 10µV 100µV 0.012+2 0.012+7 50V 100µV 1mV approx. 10M 0.016+5 0.016+10 500V 1mV 10mV 0.016+2 0.016+7 1000V 10mV 100mV The accuracy at 500mV range is specified after zero compensation through the REL operation. Max. input voltage: 500mV to 5V range ± 800V (continuous) 50V to 1000V range ± 1100V (continuous) Resolution and noise rejection Resolution Sample Rate NMRR CMRR 5.5-digit SLOW 55dB or more 120dB or more 5.5-digit MID 55dB or more 120dB or more 4.5-digit FAST 0dB 55dB or more |

| 3. CH-B DC Volt (DCV) VOAC7520 only |
| --- |
|  | Range Resolution Input Resistance Accuracy* 4.5-digit SLOW/MID FAST 5V 100µV CH-B:H to CH-B:L 10M ± 3% CH-B:H to CH-A:L 5M ± 3% CH-B:L to CH-A:L 5M ± 3% 0.025+2 0.025+30 50V 1mV 0.025+8 300V 10mV 0.025+5 Max. input voltage: ± 300V between CH-A L and CH-B ± 300V Resolution and noise rejection Resolution Sample Rate NMRR CMRR Isolation between CH-A and CH-B 4.5-digit SLOW/MID 55dB or more 120dB or more 56dB or more 4.5-digit FAST 0dB 55dB or more |

| 4. AC Volt (ACV, DC+ACV) detection of True RMS |
| --- |
|  | Range Resolution Measurement Range Input Resistance 5.5-digit SLOW MID/FAST 500mV 1µV 15Hz to 100kHz 200Hz to 100kHz less than approx. 1M 100pF 5V 10µV 50V 100µV 500V 1mV 45Hz to 100kHz 750V 10mV 45Hz to 20kHz 200Hz to 20kHz Accuracy: SLOW Sample (Sine wave) Frequency Accuracy* 15Hz to 45Hz 0.5+150 45Hz to 100Hz 0.25+150 100Hz to 30kHz 0.2+150 30kHz to 100kHz 0.5+300 Coefficient to input other than sine wave Crest Factor Frequency 1 to 1.5 1.5 to 2 2 to 3 15Hz to 30kHz 0.05% 0.15% 0.30% 30kHz to 100kHz 0.20% - - Resnponse time Sample Rate Resolution Reading Rate Response Time SLOW 5.5-digit 4 times/sec less than 3 sec MID/FAST 5.5-digit 20 times/sec less than 2 sec Max. input voltage: 780Vrms, ± 1100V DC (continuous) In the case of DC+ACV, 500 (less than 45Hz) or 300 (45Hz or higher) must be added to the value of Accuracy digit. Sample rate of FAST becomes the same values as MID (approx. 20 times/sec). |

| 5. DC Current (DCA) |
| --- |
|  | Range Resolution Accuracy* Input Resistance 5.5-digit 4.5-digit SLOW/MID FAST 5mA 10nA 100nA 0.05+7 0.05+17 less than 150 50mA 100nA 1µA less than 15 500mA 1µA 10µA less than 2 10A 100µA 1mA 0.2+7 0.2+17 less than 0.1 Auto range is not available at 5mA to 500mA range and 10A range because of using different input terminals. Max. input current: 5mA to 500mA range 500mA (FUSE 0.5A) 10A range 10A (FUSE 15A) |

| 6. AC Current (ACA, DC+ACA) |
| --- |
|  | Range Resolution Measurement Range Input Resistance 5.5-digit SLOW/MID FAST 5mA 10nA 15Hz to 5kHz 200H to 5kHz less than 150 50mA 100nA less than 15 500mA 1µA 45Hz to 5kHz less than 2 10A 100µA less than 0.1 Accuracy: SLOW Sample (Sine wave) 5% or more against the range Frequency Accuracy* Crest Factor 1 to 1.5 1.5 to 2 2 to 3 15Hz to 45Hz 1+200 0.05% 0.15% 0.30% 45Hz to 1kHz 0.4+200 1kHz to 5kHz 5.0+200 Response time Sample Rate Resolution Reading Range Response time SLOW 5.5-digit 4 times/sec less than 3 sec MID/FAST 5.5-digit 20 times/sec less than 2 sec Max. input current: 5mA to 500mA range 500mA (FUSE 0.5A) 10A range 10A (FUSE 15A) DC Component on input current must be included in the Max. input current. In the case of 10A range at 45Hz to 1kHz, 0.3 must be added to %. In the case of DC+ACA, 500 (less than 45Hz) or 300 (45Hz or higher) must be added to the value of Accuracy digit. Sample rate of FAST becomes the same value as MID (approx. 20 times/sec). |

| 7. Resistance (2 Wire/4 Wire) 4 Wire: VOAC7521A only |
| --- |
|  | Range Resolution Accuracy* Test Current SLOW/MID FAST SLOW/MID FAST 50 0.1m 1m 0.025+10 0.025+15 approx. 10mA 500 1m 10m 0.014+3 0.014+8 approx. 10mA 5k 10m 0.1 approx. 1mA 50k 0.1 1 approx. 100µA 500k 1 10 0.015+3 0.015+33 approx. 10µA 5M 10 10 0.033+30 0.033+30 approx. 1µA 50M 100 100 0.25+30 0.25+30 approx. 100nA 500M 1k 1k 1.5+50 1.5+50 approx. 10nA Max. input voltage: ± 500V peak Open circuit test voltage: less than 12V The accuracy at 50 to 5k range are specified after zero compensation through the REL operation. Sample rate of FAST at 5M to 500M range becomes the same value as MID (approx. 20 times/sec). |

| 8. Low-Power Resistance (2 Wire) |
| --- |
|  | Range Resolution Accuracy* Test Current SLOW/MID/FAST SLOW/MID FAST 500 10m 0.1+5 0.1+15 approx. 1mA 5k 0.1 approx. 100µA 50k 1 approx. 10µA 500k 10 0.2+30 0.2+40 approx. 1µA 5M 100 0.2+30 approx. 100nA 50M 1k 1.5+30 1.5+30 approx. 10nA Max. input voltage: ± 500V peak Open circuit test voltage: less than 12V The accuracy at 500 to 5k range are specified after zero compensation through the REL operation. Sample rate of FAST at 5M to 500M range becomes the same value as MID (approx. 20 times/sec). |

| 9. Diode |
| --- |
|  | Test Current Measurement Range Accuracy* Open Circuit Test Voltave Max. Input Voltave approx. 1mA or 10mA 0.1mV to 5.0999V 0.014+13 less than 12V ± 500V peak |

| 10. Temperature |
| --- |
|  | Thermo Couple Measurement Range Resolution Accuracy* Max. Input Voltave R -50 to +1768 °C 0.1°C 0.2+30 ± 500V peak K(CA) -270 to +1372 °C 0.1+15 T(CC) -270 to +400 °C 0.15+15 J(IC) -210 to +1200 °C E(CRC) -270 to +1000 °C Resolution: 4.5-digits, Sample rate at SLOW/MID/FAST: approx. 2 times/sec |

| 11. Frequency (AC couple, Crest Factor: less than 3) |
| --- |
|  | Sample Rate Reading Rate (Gate time) Display Digits and Measurement Range Accuracy* SLOW approx. 0.5 times/sec (1s) 6-digit 15.0000Hz to 1.00000MHz 0.02+2 MID approx. 4 times/sec (100ms) 5-digit 15.000Hz to 1.0000MHz FAST approx. 10 times/sec (10ms) 4-digit 150.00Hz to 1.000MHz AUTO range of ACV must be used with input attenuator. Max. input voltage: 780 Vrms, ± 1100V peak |

| 12. Chart for combination of Dual Function |
| --- |
|  | DCV CH-B DCV (*) ACV DC+ ACV DCA ACA DC+ ACA 2 Wire 4 Wire(**) Hz °C DCV Note 1 CH-B DCV (*) - ACV DC+ACV DCA ACA DC+ACA 2 Wire Note 2 4 Wire(**) - Note 2 Hz °C Note 1 : Available : have a limitation : N/A - : not provided Note 1. Thermocouple EMF and temperature are simultaneously displayed. When measuring voltage and temperature simultaneously, use CH-A on the VOAC7520 for temperature while using CH-B for DCV. Note 2. With the same measuring resistance. (*) CH-B DCV: VOAC7520 only (**) 4 Wire: VOAC7521A only |

| 13. General |
| --- |
|  | MATH Moving Average, Scale, Decibel (dBm, dBµ), Difference, Statistics (MAX, MIN, , ), Comparison (COMP), Arithmetic Calculation between Dual Function Memory DATA Max. 3000 data with 10 msec resolution time mark (Elapsed time) SET UP 10 Interfaces (Full Remote) Standard RS-232 Option LAN, GP-IB Power Supply Voltage 100V, 110V, 220V, 240V Frequency 50Hz, 60Hz Power Consumption less than 21 VA (includes options) Operation Temperature and Humidity 0°C to 50°C (less than 80%RH) no condensation, less than 70%RH at 40°C to 50°C Storage Temperature and Humidity -20°C to 60°C (less than 70%RH) no condensation, includes operation temperature Size Measurement 210W x 99H x 353L mm (Options are built into the main unit) Weight less than 3.5kg (includes options) Standard Accessories Fuse, Test Leads, Alignment Screwdriver, Operation Manual, Power cord |